## data
Misc Notes / Thoughts
Scanner is a very new system with an AI recon for Adult Abd pelvis and scanner previously set to pick mAs based on expectation of AI recon. Scanner currently has very few adult patient data points and no pediatric
this system had very low patient doses. Just did a followup test of new protocols for ACR submission and it's CNR will not pass. Currently going to try to collect some data with Sure Exposure 3D set to Standard instead of AiCE Standard.
a few studies might have had their mAs limited because of the limits for the min and max mA….. However this would not impact their assessment of Dw.
Dw for 32 cm with inner 16 and 10 cm both removed a pretty good match for a 70 Kg patient.. (phantom = 29.4 cm, low statistic patient average of 28.6 cm)
Dw for 32 cm with inner 16 removed but 10 cm resting inside larger hole pretty good match for a 84 Kg patient.. (phantom = 31.2 cm, low statistic patient average of 31.5cm)
Dw for 16cm CTDI a little low relative to adult head patients so far (phantom = 16.9 cm, low statistic patient average of 18.9 cm). The mA matched really well when the 16 cm was scanned on patient table and not head holder.
Dw for ACR phantom on the uniformity slice, not in its positioning cradle, is around 20 cm. I a little high for an adult head. However I wonder if a simple (or weighted average) of the mA a scanner would select for it and for the 16 cm CTDI could be used for a adult head?
A similar approach might be feasible for ped abd and ped head… but I need more data on what an average Dw would be for those patient populations. The Dw for the ped/abd should be close to that for an adult head.
in future… if doing more sampling… make sure to set min and max mA limits to lowest and highest from the beginning of data acquisition
ROI sizes didn't change much as I didn't have to redraw the ROI unless changing from a different phantom outer diameter.
I put a about 1 cm of foam underneath the CTDI phantom. This was helpful as it allowed easy/quick use of circular ROI that could enclose the CTDI phantom without inclusion of patient table.
PATIENT DATA
category | patient sample size | average mA | average mAs | Average effective mAs | Average AP (cm) | Average Lat (cm) | Average Dw (cm) | Notes
Adult Abd - mid liver - ~70kg | 6 | 92.8333 | 46.4167 | 57.0931 | 24.6167 | 31.5833 | 28.5631
Adult Abd - mid liver - avg cdc weight ~ 84 kg | 4 | 125 | 62.5 | 76.8758 | 27.025 | 34.75 | 31.5132
Adult Head | 5 | 338 | 253.5 | 405.6 | 19.32 | 15.36 | 18.9131 | mA likely saturated on technique all but one patient was at 340 mA, other was 330
PHANTOM TEST DATA | [7 empty] | Theoretical Estimation Based on acrylic volume, 124 HU for acrylic, and -1000 for air |  |  |  | Empirical Data
Phantom Configuration | Scanner Used | Date | Acquisition Parameters | kVp | mA | mAs | effective mAs | area of acrylic in phantom ROI (cm^2) | area of phantom ROI (acrylic and internal air) (cm^2) | Estimated Average HU | ~ Dw (cm) | Drawn ROI (excludes support surface, may include extra air) (cm^2) | Average HU for ROI | Calculated Dw (cm) | notes
32cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 120 | 164 | 82 | 100.861 | 804.2477 | 804.2477 | 124 | 33.926 | 855.8 | 59.75 | 33.9815
32cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 135 | 114 | 57 | 70.1107 | 804.2477 | 804.2477 | 124 | 33.926 | 855.8 | 64.38 | 34.0557
32cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 100 | 312 | 156 | 191.8819 | 804.2477 | 804.2477 | 124 | 33.926 | 855.8 | 54.35 | 33.8948
32cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 80 | 580 | 290 | 356.7036 | 804.2477 | 804.2477 | 124 | 33.926 | 855.8 | 41.83 | 33.693
32cm outer + 16 cm - 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 120 | 124 | 62 | 76.2608 | 725.7079 | 804.2477 | 14.2344 | 32.2269 | 855.8 | -43.82 | 32.2783
32cm outer + 16 cm - 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 135 | 90 | 45 | 55.3506 | 725.7079 | 804.2477 | 14.2344 | 32.2269 | 855.8 | -39.96 | 32.3434
32cm outer + 16 cm - 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 100 | 230 | 115 | 141.4514 | 725.7079 | 804.2477 | 14.2344 | 32.2269 | 855.8 | -48.79 | 32.1943
32cm outer + 16 cm - 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 80 | 544 | 272 | 334.5633 | 725.7079 | 804.2477 | 14.2344 | 32.2269 | 855.8 | -60.28 | 31.9993
32cm outer - 16 cm - 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 120 | 90 | 45 | 55.3506 | 603.1858 | 804.2477 | -157 | 29.3808 | 855.8 | -205.93 | 29.4151
32cm outer - 16 cm - 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 135 | 80 | 40 | 49.2005 | 603.1858 | 804.2477 | -157 | 29.3808 | 855.8 | -202.64 | 29.476 | likely throttled by min mA
32cm outer - 16 cm - 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 100 | 146 | 73 | 89.7909 | 603.1858 | 804.2477 | -157 | 29.3808 | 855.8 | -209.93 | 29.3409
32cm outer - 16 cm - 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 80 | 362 | 181 | 222.6322 | 603.1858 | 804.2477 | -157 | 29.3808 | 855.8 | -209.87 | 29.3421
32cm outer - 16 cm + 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 120 | 104 | 52 | 63.9606 | 681.7256 | 804.2477 | -47.2344 | 31.2351 | 855.8 | -102.44 | 31.2732
32cm outer - 16 cm + 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 135 | 80 | 40 | 49.2005 | 681.7256 | 804.2477 | -47.2344 | 31.2351 | 855.8 | -98.28 | 31.3456 | likely throttled by min mA
32cm outer - 16 cm + 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 100 | 194 | 97 | 119.3112 | 681.7256 | 804.2477 | -47.2344 | 31.2351 | 855.8 | -106.48 | 31.2028
32cm outer - 16 cm + 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of 0.813, rot time of 0.5s, Sure Exp 3D AiCE Standard | 80 | 480 | 240 | 295.203 | 681.7256 | 804.2477 | -47.2344 | 31.2351 | 855.8 | -117.91 | 31.0026
16 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of .813, rot time of 0.5s, Sure Exp 3D AiCE Standard, min mA reduced to 10 | 120 | 14 | 7 | 8.6101 | 201.0619 | 201.0619 | 124 | 16.963 | 224 | 6.81 | 16.9454 | was throttling at min mA of 100 yielding effective mAs at 40, manually changed min mA to 10.
16 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of .813, rot time of 0.5s, Sure Exp 3D AiCE Standard, min mA reduced to 10 | 135 | 10 | 5 | 6.1501 | 201.0619 | 201.0619 | 124 | 16.963 | 224 | 11.06 | 16.9812
16 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of .813, rot time of 0.5s, Sure Exp 3D AiCE Standard, min mA reduced to 10 | 100 | 30 | 15 | 18.4502 | 201.0619 | 201.0619 | 124 | 16.963 | 224 | 1.55 | 16.9011
16 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of .813, rot time of 0.5s, Sure Exp 3D AiCE Standard, min mA reduced to 10 | 80 | 60 | 30 | 36.9004 | 201.0619 | 201.0619 | 124 | 16.963 | 224 | 11.51 | 16.9849
16 cm - 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of .813, rot time of 0.5s, Sure Exp 3D AiCE Standard, min mA reduced to 10 | 120 | 10 | 5 | 6.1501 | 122.5221 | 201.0619 | -315.0625 | 13.2418 | 224 | -389.55 | 13.1948 | likely throttling at min mA of min mA settings
16 cm - 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of .813, rot time of 0.5s, Sure Exp 3D AiCE Standard, min mA reduced to 10 | 135 | 10 | 5 | 6.1501 | 122.5221 | 201.0619 | -315.0625 | 13.2418 | 224 | -387.06 | 13.2217 | likely throttling at min mA of min mA settings
16 cm - 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of .813, rot time of 0.5s, Sure Exp 3D AiCE Standard, min mA reduced to 10 | 100 | 14 | 7 | 8.6101 | 122.5221 | 201.0619 | -315.0625 | 13.2418 | 224 | -392.57 | 13.1622
16 cm - 10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of .813, rot time of 0.5s, Sure Exp 3D AiCE Standard, min mA reduced to 10 | 80 | 34 | 17 | 20.9102 | 122.5221 | 201.0619 | -315.0625 | 13.2418 | 224 | -401 | 13.0705
10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of .813, rot time of 0.5s, Sure Exp 3D AiCE Standard, min mA reduced to 10 | 120 | 10 | 5 | 6.1501 | 78.5398 | 78.5398 | 124 | 10.6019 | 89.63 | -20.87 | 10.5707 | likely throttling at min mA of min mA settings
10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of .813, rot time of 0.5s, Sure Exp 3D AiCE Standard, min mA reduced to 10 | 135 | 10 | 5 | 6.1501 | 78.5398 | 78.5398 | 124 | 10.6019 | 89.63 | -15.83 | 10.5978 | likely throttling at min mA of min mA settings
10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of .813, rot time of 0.5s, Sure Exp 3D AiCE Standard, min mA reduced to 10 | 100 | 14 | 7 | 8.6101 | 78.5398 | 78.5398 | 124 | 10.6019 | 89.63 | -23.83 | 10.5547 | likely throttling at min mA of min mA settings
10 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Abd/Pel WO, pitch of .813, rot time of 0.5s, Sure Exp 3D AiCE Standard, min mA reduced to 10 | 80 | 30 | 15 | 18.4502 | 78.5398 | 78.5398 | 124 | 10.6019 | 89.63 | -36.82 | 10.4842 | likely throttling at min mA of min mA settings
16 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Adult Head - phantom on table, (Brain WO), Sure Exp 3D Standard, pitch 0.625, rot time = 0.75s | 120 | 340 | 255 | 408 | 201.0619 | 201.0619 | 124 | 16.963 | 222 | 9.16 | 16.8893 | CTDIvol = 51.1
16 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Ped Abd (Ab Chil 16-30kg), Sure Exp 3D 16-130 kg, pitch 0.813, rot time = 0.5 s | 100 | 58 | 29 | 35.6704 | 201.0619 | 201.0619 | 124 | 16.963 | 222 | 2.15 | 16.8305 | CTDIvol = 0.6
16 cm (holes filled) | NGPT, Canon Aquilion Prime SP | 2026-09-14 | Adult Head - phantom in head holder, (Brain WO), Sure Exp 3D Standard, pitch 0.625, rot time = 0.75s | 120 | 289.3333 | 217 | 347.2 | 201.0619 | 201.0619 | 124 | 16.963 | 207 | 83.93 | 16.9021 | CTDIvol = 43.7
ACR Phantom, uniformity/distance slice, no positioning/support cradle | MDC, Canon Aquilion Prime SP | 2026-04-27 |  | 120 | [5 empty] | #DIV/0! | #DIV/0! | 335.6 | -63.6 | 20.0031
ACR Phantom, uniformity/distance slice, no positioning/support cradle | MDC, Canon Aquilion Prime SP | 2026-04-27 |  | 135 | [5 empty] | #DIV/0! | #DIV/0! | 329.3 | -48.74 | 19.971
ACR Phantom, uniformity/distance slice, no positioning/support cradle | MDC, Canon Aquilion Prime SP | 2026-04-27 |  | 100 | [5 empty] | #DIV/0! | #DIV/0! | 332.4 | -53.52 | 20.0143
ACR Phantom, uniformity/distance slice, no positioning/support cradle | MDC, Canon Aquilion Prime SP | 2026-04-27 |  | 80 | [5 empty] | #DIV/0! | #DIV/0! | 318.7 | -14.93 | 19.9931
Other Theoretical Test Configurations (with probe holes removed) - NOT VALIDATED ON A CT SCANNER
Phantom Configuration | area of acrylic in phantom ROI (cm^2) | area of phantom ROI (acrylic and internal air) (cm^2) | Estimated Average HU | ~ Dw (cm)
32 cm outer ring only, 4 holes empty | 597.7858 | 804.2477 | -164.5469 | 29.249
16 cm ctdi, w/ 10 cm, 5 holes empty | 194.3119 | 201.0619 | 86.2654 | 16.6758
16 cm ctdi, w/ 10 cm, 9 holes empty | 188.9119 | 201.0619 | 56.0776 | 16.4425
16 cm minus 10 cm, 4 holes empty | 117.1221 | 201.0619 | -345.2502 | 12.9467
10 cm ctdi, 5 holes empty | 71.7898 | 78.5398 | 27.3993 | 10.1361
5 fill rods rubber banded together | 6.75 | 15 | -494.2 | 3.1081
9 fill rods rubber banded together | 12.15 | 25 | -453.736 | 4.1699
13 fill rods rubber banded together | 17.55 | 35 | -436.3943 | 5.0116